## Slide 1: How to build an NMR spectrometer?
Biological NMR
www.bionmr.ethz.ch
Takuya Segawa
BioNMR group seminar
November 3, 2022
… or NMR for everyone

## Slide 2
The NMR world map
China
United States
Germany
United Kingdom
India
France
Japan
Italy
Russian Federation
Top 10
Citations in Scopus (2021), keyword = NMR, by Maria Pechlaner (Infozentrum)

## Slide 3
The NMR world map (per capita)
Switzerland
Denmark
Slovenia
Sweden
Finland
Austria
Iceland
Luxembourg
Netherlands
Top 10
Citations in Scopus (2021), keyword = NMR, by Maria Pechlaner (Infozentrum)

## Slide 4: Inaccessibility of NMR in the «Global South»
To change this, we need…
…a cheap NMR spectrometer
…a cryogen-free NMR spectrometer
…an open-source NMR spectrometer / community
Roland:
«CHF 10’000 NMR spectrometer»
photo-CIDNP signal enhancement
Smartphone App controlled
for education and research
No NMR, because of…
…prohibitive procurement costs
…high operating costs
…inaccessibility of liquid Helium
Without NMR spectroscopy…
…no chemical structure determination
…no «hands-on» quantum physics experiment
…exclusion from 75-years-old MR technology

## Slide 5: Building an NMR spectrometer: The ingredients
Magnet
Console
Preamp
Computer
Probe

## Slide 6: Building an NMR spectrometer: The ingredients
Magnet
Field strength :
Sensitivity (SNR / time):
Resolution: (chemical shift range / J-couplings):
Field homogeneity:
goal: resolution of ca. 1 Hz
  0.1 ppm for 10 MHz
  1 ppb for 1 GHz!
Shim coils
 Improvement of up to 103 in homogeneity
Probe
RF coil for transmission and reception with LC circuit for tuning and matching  for every channel
Console
RF frequency, RF pulse, amplifier, RF detection, digitizer, clock, etc.  for every channel
Preamplifier
low-noise amplifier & transmitter/receiver switch
Computer
 Data: processing and display

## Slide 7: Benchtop NMR – The market
Magritek, Spinsolve 90
Spinsolve series on the market since 2012
CHF ?
Broadband probe: 1H, 19F, 13C, 31P, 7Li, 11B, 29Si, 129Xe, 2D, 15N
3 D gradients + gradients for DOSY (optional)
115 kg
1H sensitivity (1% Ethyl Benzene): ≥ 240:1
Bruker, Fourier 80
CHF 100’000 – 150’000
Topspin software
External lock on 19F
option: gradients, temperature (25 – 60 °C)
100 kg
1H sensitivity (1% Ethyl Benzene): ≥ 240:1

## Slide 8: Benchtop NMR – The market
QMagnetics, 125 MHz
Highest-field benchtop on the market
< USD 40’000?
Software: open-source (in development?)
No lock («soft lock»)
28 kg
Active volume: only 1 mL  flow system
1H sensitivity (1% Ethyl Benzene): ≥ 14:1
Nanalysis, 100 MHz
CHF ?
Topspin software
lock on 2H
110 kg
1H sensitivity (1% Ethyl Benzene): >220 : 1

## Slide 9: Photo-CIDNP at low field @ Fourier 80
with Barbara Czarniecki (Bruker, Fällanden)
1H 80 MHz NMR
Molecule: 100 mM tryptophan
Dye: 25 mM fluorescein
O2 quencher:
  enzyme system glucose oxidase / catalase & 5 mM glucose
Solvent: 100 % H2O
Pulse sequence: zgcppr (a composite pre-saturation pulse)
Laser irradiation: 450 nm, ca. 1 W, optical fiber into NMR tube
Number of scans: 4
laser off
laser on
5 mM glucose
0.1 mM tryptophan
(aromatic protons)
0.1 mM tryptophan
(aliphatic protons)

## Slide 10: Pioneers in open-source MR: 1) Kazuyuki Takeda’s Opencore NMR (2007 – now)
K. Takeda, Highly Customized NMR Systems Using an Open-Resource, Home-Built Spectrometer, Annu. Rep. NMR Spectrosc., 74, (2001)
https://doi.org/10.1016/B978-0-08-097072-1.00007-8
Core part:
Field-Programmable Gate Array (FPGA)
“Hardware becomes Software”
Performs all digital job:
pulse programmer
direct digital synthesizers
digital receiver
digital low-pass filter
PC interfaces for command/data transfer
Costs:
ca. 5000 €
Frequency range
up to 600 MHz (14 T for 1H NMR)
Use same data format  use processing software

## Slide 11: Pioneers in open-source MR : 2) Open-Source MRI
V. Negnevitsky et al. (2022), arXiv:2208.01616 | 

## Slide 12: Pioneers: 2) Open-Source MRI
T. Guallart-Naval, J. Alonso et al., Sci. Rep., 12:13147 (2022) https://doi.org/10.1038/s41598-022-17472-w
T. Guallart-Naval, J. Alonso et al. (2022) arXiv:2203.11314
= 70 mT MRI
< EUR 50’000
Mobile (250 kg)
in 10 min

## Slide 13: Pioneers: 2) Open-Source MRI
Key part: SDR Lab from Red Pitaya (Slovenia)
FPGA: Xilinx Zynq 7020
2 RF inputs (16-bit resolution)
2 RF outputs (14-bit resolution)
Max. RF: 122.88 MHz
  Up to ca. 60 MHz spectrometer (1H ca. 1.5 T)
Software on an SD card (exchangeable)
EUR 500
EUR 625
SDR: “software-defined radio”

## Slide 14: MaRCoS – “Red Pitaya”-based console
 let’s use this as our NMR console!
ü

## Slide 15: Red Pitaya @Riek Lab
Generation of RF pulses with defined RF:
   Measured on oscilloscope
Detection of RF signal, sent by RF signal generator:
  5.01 MHz detected with a carrier frequency of 5 MHz
  10 kHz signal
  24.46 MHz for 80 ms

## Slide 16: The Magnet: 0.6 T for 25 MHz 1H NMR
Field strength:
0.6 T  25 MHz
Homogeneous volume:
100 mL / 5 mm tube
Homogeneity:
0.1 ppm FWHM  2.5 Hz
Magnet material:
temperature-compensated Samarium-Cobalt  50 ppm/K
Weight:
ca. 5 kg
Price for prototype
USD 26,532.00 + VAT
Price for one magnet (N = 10)
USD 8,500.00 + VAT
Expected delivery:
end of 2022/
beginning of 2023
Passive and active shim assembly
By SABR Enterprises, LLC (USA)
ü
Solenoid coil for RF excitation/detection
ü
Ca. CHF 1000 left to build rest of spectrometer!

## Slide 17: Shimming
ü
67 MHz 1H NMR
without shim
with shim
A. McDowell and M. Conradi, J. Magn. Reson., 281, 7 (2017)
A.F. McDowell, Magnetic Field Correction System, US Patent 9285441, 2016.
Challenges:
No existing algorithm for «Auto-Shim»
Need of current supply (32 channels)

## Slide 18: Preamp and Switch
Transmit/Receive switch (duplexer)
from Mini-Circuits:
Low-noise amplifier
from Mini-Circuits:
EUR 125
Noise Figure: 2.9 dB
Frequency: 0.1-500 MHz
Albannay (Bannay) Mohammed
Alexander Däpp
ü

## Slide 19: And what about photo-CIDNP?
Source: NZZ (31.10.2022)
Use high-power LEDs from Indoor Farming

## Slide 20: Excite photo-CIDNP with LED
High Power LEDs –
CREE Ltd
XP-E Blue (455 nm), 45.7 lm
CHF 1.60

## Slide 21
Challenge of low-field resolution
Simulated/Predicted with MNova

## Slide 22: Chances in Sensitivity Enhancement through photo-CIDNP
Field strength :
Sensitivity (SNR / time):
For Hyperpolarization:
 , where = Polarization
Field dependence of photo-CIDNP
Different Molecules
Red: HOPI
Green: tryptophan
A) AT 12 Different Dyes B) fluorescein
F. Torres, A. Sobol, J. Greenwald, A. Renn, O. Morozova, A. Yurkovskaya, R. Riek , PCCP, 23, 6441 (2021)
.
 Optimize Dye/Molecule Pair
 Maximize Polarization
Notes: Fig. 5 Field dependent photo-CIDNP polarization calculated from the HFCC of HOPI (red) and tryptophan (green) in the presence of AT12 (A), or
fluorescein (B). The g-value of AT12 is unknown and the curves were simulated for different potential values: 2.0036 (solid), 2.0037 (dashed), 2.0038 (dotdashed),
2.0039 (dotted), and 2.0040 (loosely dotted).

## Slide 23: To Do
Software / Smartphone App
Temperature insulation/control of magnet  field stability
Automatization: tuning/matching, shimming, locking on water (1H2O), pulse calibration, water suppression (pre-saturation)
Spectral simplification: decoupling, ML approaches to extract more information?
Move to 2nd spectrometer prototype with different magnet design
Applications?
Research: photo-CIDNP drug screening  record intensity changes  spectral resolution not crucial!
 competitive binding with low-field optimized dye/molecule pair
Teaching: «normal» NMR with highly concentrated samples (100 mM)
  Decoupling, relaxation measurements
For advanced NMR users
For non-experienced NMR users
Start a collaboration with a university of the «Global South»

## Slide 24: Roland
Thank you!

## Slide 25: The photo-CIDNP group
Thank you!

## Slide 26
Albannay (Bannay) Mohammed
Quad Systems
Mazin Jouda
KIT
Jan Korvink
KIT
Vlad Negnevitsky
Oxford Ionics/MaRCoS
Alexander Däpp
LPC electronics
Tiago Ferreira
LPC electronics
Thank you!

## Slide 27: Switzerland: The NMR world power
“A new fascination for the NMR experiment”
I remember, in the winter of our first experiments […] looking on snow with new eyes. There the snow lay around my doorstep - great heaps of protons quietly precessing in the earth’s magnetic field.
Purcell, Nobel Lecture, 1952
https://www.nobelprize.org/uploads/2018/06/purcell-lecture.pdf